## Slide 1
NON-STERILE PREP
STERILE PREP
DEVICE CONTROLS
SOFTWARE
TROUBLESHOOTING
POST PROCEDURE
Sonata System for Sonography-Guided Transcervical Fibroid Ablation
QUICK REFERENCE GUIDE
Refer to the Sonata System Manual for complete operating instructions.

## Slide 2
QUICK REFERENCE GUIDE | NON-STERILE PREP
1
2
3
6
5
4
Lower System Cart to move into position.
Connect power supply.
Adjust cart height for viewing.
Lock at least 2 wheels.
continued >>
NON-STERILE PREP
Turn on using switch on back of RF Generator.
Position System Cart and Lock 2 Wheels
Open Ultrasound Console (laptop)
Power ON Ultrasound Console with Button
Connect Footswitch
Power ON RF Generator
Enter Patient Data
Login to Ultrasound Console.
Enter Patient Data to add to digital file.
Click “Save and close” button.
Firmly connect pneumatic tube of Footswitch to RF Generator.

## Slide 3
QUICK REFERENCE GUIDE | NON-STERILE PREP
7
8
9
Check expiration date on kit and contents.
Remove both non-sterile Dispersive Electrodes from pouches.
Apply 1 per thigh.
Must have full contact with skin.
Direct cable toward knee.
Indicators blink when not properly connected.
Indicators will disappear from the display when all components are properly connected.
10
11
12
NON-STERILE PREP
Open Sonata Case Kit
Apply Dispersive Electrodes
Connect Dispersive Electrode Cables
Confirm Connections
Connect and Lock IUUS Probe to Ultrasound Console
Connect RFA Handpiece to RF Generator

## Slide 4
STERILE PREP
QUICK REFERENCE GUIDE | STERILE PREP
1
2
3
6
5
4
250 cc of Hypotonic fluid
2 sterile 60 cc syringes
Standard D&C Kit with dilators to 27 Fr (9mm)
New RFA Handpiece
Sterile IUUS Probe
Use sterile technique to remove RFA Handpiece from packaging.
Grasp from center to protect shaft.
Place sterile IUUS Probe on draped table.
Connect tips of IUUS Probe to RFA Handpiece and snap handles together.
RFA Handpiece
Confirm that all controls are working.
Deploy Needle Electrodes and retract until Articulating Tip is back to 0°.
Remove air bubbles from syringe and connect to infusion port.
Flush Treatment Device with hypotonic fluid.
60cc syringe
Prepare Draped Table with Required Supplies
Open RFA Handpiece
Open IUUS Probe
Assemble Treatment Device
Inspect Treatment Device
Flush Treatment Device
IUUS Probe
Verify tips are connected.
Return Articulating Tip to 0°.
Needle Electrodes should be straight.

## Slide 5
SMART Guide™| Setting Margins of Ablation in Real Time
DEVICE CONTROLS
Ablation Zone
IUUS Probe Tip
Thermal Safety Border
Needle Electrode Origin
Introducer Tip Marker
Needle Electrode Tip Markers
Introducer Guide
The IUUS Probe Tip illustration at the top of the screen represents Tip angle (45° or 60°).
The Introducer Guide is a blue dashed line that indicates the path of Introducer deployment. -----
The Ablation Zone is indicated by the red inner ellipse and controlled with the Ablation Guide Knob.
The Thermal Safety Border is indicated by the green outer ellipse and controlled with the Ablation Guide Knob.
The Introducer Tip Marker is shaped like a yellow diamond  and the Needle Electrode Origin is shaped like a yellow star. 
The Needle Electrode Tip Markers are the 3 yellow lines that represent where the Needle Electrode tips will be positioned when deployed. The yellow lines turn red when the Needle Electrodes are fully deployed. – –



## Slide 6
QUICK REFERENCE GUIDE | DEPLOY 1-2-3-4
5
Depress, Slide, Rotate Ablation Guide Knob
Advance Introducer
Deploy Needle Electrodes
Depress Articulation Lever to Angle Tip
1
2
3
4

## Slide 7
QUICK REFERENCE GUIDE | RETRACT 4-3-2-1
Articulate Lever to Return Tip to 0°
Retract Needle Electrodes
4
3
2
1
Retract Introducer
Depress and Return Ablation Guide Knob

## Slide 8
Remove Footswitch.
Wipe with medical grade low alcohol disinfectant.
Store on System Cart.
QUICK REFERENCE GUIDE | POST PROCEDURE
1
2
3
6
5
4
continued >>
Inspect skin surface for burns.
RFA Handpiece
IUUS Probe
Disconnect Dispersive Electrodes
Remove Dispersive Electrodes & Check Skin
Disconnect, Clean, & Store Footswitch
Disassemble Treatment Device
Disconnect RFA Handpiece & IUUS Probe
Slide RFA Handpiece forward to separate.
Lift release latches simultaneously.
Manage Video Recording
Click “Keep” to store video on hard drive.
Click “Discard” to permanently delete video.
POST PROCEDURE

## Slide 9
12
11
10
QUICK REFERENCE GUIDE | POST PROCEDURE
7
8
9
Dispose in Medical Hazard Sharps container.
Wipe down cart with medical grade low alcohol disinfectant.
Unplug power cord.
Unlock wheels.
Lower completely prior to transport.
Use back handle to move.
Turn off using switch on back of RF Generator.
Send IUUS Probe for Reprocessing
Dispose of RFA Handpiece
Power OFF Ultrasound Console with Power Icon
Power OFF RF Generator
Clean System Cart with Low Alcohol Disinfectant & Store
Pre-Clean IUUS Probe
Carefully remove visible soil from IUUS Probe imaging surface, shaft, handle, and cable.
Place in Reprocessing Tray to protect IUUS Probe Tip.

## Slide 10
Problem | Troubleshooting | 
Unable to Enter New Patient Data | Click on the previously entered patient name. Choose “New Patient” button. | 
QUICK REFERENCE GUIDE | TROUBLESHOOTING
Ultrasound Console Connection to RF Generator
Dispersive Electrode Connections to RF Generator
RFA Handpiece Connection to RF Generator
If a Component Connection Indicator is blinking, check the connections of the component(s) displayed.
Component Connection Indicators are Blinking
TROUBLESHOOTING

## Slide 11
QUICK REFERENCE GUIDE | TROUBLESHOOTING
Problem | Troubleshooting
When attempting to “Click on Tip” the arrow does not change to a “+”. (User hears an audible tone when clicking.) | The Introducer Tip is out of range. Retract Introducer Tip to align with Introducer Tip Marker .
 | Move System Cart to a different location. Change to a different wall outlet (isolated from other equipment). Make sure there is contact between tissue and IUUS Probe. Check the IUUS Probe cable is not wrapped around the System Cart or other electrical equipment.
+
Electromagnetic Interference (EMI)


## Slide 12
QUICK REFERENCE GUIDE | SOFTWARE CONTROLS
SOFTWARE
Ultrasound Imaging Presets
Time and Date
Patient Data
RF Generator Status
IUUS Probe Tip
Preferences
Power OFF
Image Review
Still Image Capture
Freeze or Unfreeze
Articulation Angle
Component Connection
Indicators
Ultrasound
Data Field

## Slide 13
ICON/IMAGE | LABEL | HOW TO USE
 | Entering Patient Data for New Patients | The Patient Data window is displayed upon login, or following a procedure. To enter a new patient: Click on the patient name (e.g. Jane Doe). Click on the “New Patient” button. Enter patient data.
 | Component Connection Indicators | Components that are not connected will blink. Components that are connected will be solid. When all components are connected, the indicators will disappear.
 | Ultrasound Imaging Presets | The default Ultrasound Image Preset is for an average sized fibroid (the middle blue circle on the left side of the screen). The presets are programmed to adjust the image depth, ultrasound frequency, and other imaging parameters. To change to a preset, click on a circle (small, medium, or large).
QUICK REFERENCE GUIDE | SOFTWARE CONTROLS
SOFTWARE
QUICK REFERENCE GUIDE | SOFTWARE CONTROLS

## Slide 14
QUICK REFERENCE GUIDE | SOFTWARE CONTROLS
ICON/IMAGE | LABEL | HOW TO USE
 | Freeze/Unfreeze | To freeze image, click on Freeze icon or the spacebar. Freeze Icon will remain gray while image is frozen. To unfreeze image, click on Freeze Icon or the spacebar.
 | IUUS Probe Articulation Angle | Articulation angle is displayed at the top of the ultrasound data field. The angles available are 0°, 45°, and 60°.
 | RF Status | Offline – RF not connected Standby – During targeting phase Ready – Ready to be activated by operator Active – Activated, heats Needle Electrodes to 105°C
 | Still Image Capture | To capture a still image, click on Still Image Capture icon.
 | Image Review | To review captured images, click on Image Review icon.

## Slide 15
QUICK REFERENCE GUIDE | MANAGE FILES
1
2
3
6
5
4
Window displays after clicking on the patient name or “Patient Data.”
Remove USB Drive and make sure RF Generator cable is reconnected.
Insert USB Drive Only in Top USB Port
Select “Manage Patient Data” Option
Reconnect RF Generator Cable
Disconnect RF Generator Cable from Top USB Port
Select Patient File(s) to be copied.
Click “Export to USB” button.
Choose Files to Export to USB
Click “Copy Selected Patient Files”
Always use a new USB drive or scan for viruses before inserting.

## Slide 16
Gynesonics, Inc. 301 Galveston Drive Redwood City, CA 94063 Telephone: +1-650-216-3860 www.gynesonics.com | Gynesonics Netherlands BV Kapershof 46 6641 JS Beuningen The Netherlands Phone: +31 24 6750555
The Sonata System is intended for diagnostic intrauterine imaging, and transcervical treatment of symptomatic uterine fibroids, including those associated with heavy menstrual bleeding. For full prescribing information, contraindications, patient selection criteria, and safety information, refer to the Sonata System Manual or visit www.gynesonics.com/sonata-system. SMART is a trademark and Sonata is a registered trademark of Gynesonics, Inc. © 2017 Gynesonics, Inc. LS 05309 Rev 02 September 2017 | 
0086